Common commands and macros with the more important ones in bold-face

(older/longer names given in parentheses first)

| a | (acqu) move to the acquisition window |
| --- | --- |
| aq # | input the acquisition time of the fid, if value not given (e.g. only aq entered) user is prompted, normally td or sw is altered and aq is left dependent |
| ad | (ased) list all experimental parameters pertinent to the current experiment for the loaded pulse program |
| apk | run automatic phase correction |
| ased | list all experimental parameters pertinent to the current experiment for the loaded pulse program |
| atma | run automatic tune and match procedure for BBI probes |
| atmm | run manual tune and match procedure for BBI probes |
| autof | turns off the autoshim (macro to call tune file of same name) |
| auton | turns on the autoshim (macro to call tune file of same name) |
| AUTOSHIM ON/OFF | turns on/off the autoshim function, only works when given in a tune file and letters capitalised |
| basl | window for setting baseline points or applying function for baseline correction (if setting user-defined baseline points, apply by sab command) |
| bk | (absd) perform automatic correction of the baseline |
| c | decrement the experiment number |
| cf | go into configuration menu |
| dir | provide a directory listing of the experiments for the current user (names & wildcards can be used) |
| del | delete data sets (names & wildcards can be used) |
| dela | delete acquisition files (names & wildcards can be used) |
| delp | delete processed files (names & wildcards can be used) |
| diro | provide a directory listing of the users (names & wildcards can be used) |
| dpa | display the values of the parameters actually used to acquire the displayed spectrum (as opposed values changed subsequently) |
| do | xf2 + abs2 + dosy2d |
| d1 # | input the post-acquisition time of the fid (to allow for relaxation in addition to that occurring during acquisition, i.e. total relaxation time for non-PT experiments = aq + d1), if value not given (e.g. only d1 entered) user is prompted, default is seconds, use m for milliseconds and u for microseconds |
| e | (edc, new) call up an existing data set or create a new data set, i.e. access/create new user, experiment no. (parameters are copied from the experiment from where the move was made if new experiment created) or processing no. |
| eda | list all experimental parameters |
| edasp | show the routing for the rf pulses |
| edau | edit automation program file, no input brings up list (names & wildcards can be used) |
| edc | call up an existing data set or create a new data set or user |
| edhead | set the probehead |
| edlock | edit the table of solvent lock parameters including field |
| edmac | edit macro, no input brings up list (names & wildcards can be used) |
| edmisc | edit miscellaneous files, e.g. integration list, peak list, etc. |
| edo | edit output device parameters |
| edp | list all processing parameters |
| edprosol | edit/view the pulse widths and power level settings |
| edsolv | edit/view the solvent table |
| edte | display the temperature control window |
| edtune | edit tune file, no input brings up list (names & wildcards can be used) |
| efp | em (exponential multiplication) + ft (Fourier transformation) + pk (phase correction using set values) |
| ej | eject the sample |
| ep | (edcpul) edit the current pulse program (edpul will list all pulse program for selection prior to editing) |
| em | exponential multiplication |
| expt | request the total time that the experiment requires |
| e2 | (edc2) for setting spectra for the dual function (display two spectra together) or for the display on the f2, f1 dimensions of 2D spectra |
| f | (efp) apply exponential multiplication to the fid, Fourier transform the fid, and apply the phase correction parameters currently in store to the spectrum |
| fid | display FID (instead of processed spectrum) |
| fidtoser | combine 1D files into serial file |
| frq | display spectrum instead of FID in acquisition window under gs |
| ft | Fourier transformation |
| fp | Fourier transform the fid (no window function applied to the fid prior to transformation) and apply the phase correction parameters currently in store to the spectrum |
| g | (go + a + efp) start the experiment, move to the acquisition window and then when finished perform efp |
| get | (getprosol) extract the pulse width values and power levels for the resident probe from the prosol table and substitute them into the appropriate parameters for the current experiment |
| go | resume a halted acquisition, data is added to what was already accumulated |
| gs | continuously acquire 1 scan and display the result without signal averaging (i.e. only the result of 1 scan is displayed = real time acquisition) |
| gsm | (gradshimau) run gradient shimming automatically twice |
| g1 | (gradshimau) run gradient shimming automatically once |
| g2 | (gradshimau) run gradient shimming automatically twice |
| g3 | (gradshimau) run gradient shimming automatically thrice |
| h | (humpcal) calculate the linewidth at half-height, 0.55%, and 0.11% |
| halt | halts the current acquisition and saves the data, if more acquisitions are to follow, they will run |
| history | the file “history” lists the commands entered at the command line and resides in the directory: C:\Bruker\XWIN-NMR\prog\curdir\bruker500 |
| i | increment the experiment number, if it does not already exist the parameters are copied from the preceding experiment |
| ii | initialise experiment interface, good to do when changes made to frequencies, pulse sequences, routing, etc., can alleviate/solve problems |
| ij | insert the sample |
| it | (atmm a) start the manual tune for the inverse probe |
| kill | request the list of processes that can be stopped from which one can be selected, killing an acquisition is the same as stop, i.e. stops the acquisition is stopped but the data is not saved N.B: if no selection required, choose cancel, closing the window by the X button will kill the selected process ! |
| l | (lockdisp) display the lock signal window |
| la | lock on acetone (lc, for CDCl3, ld for DMSO, lm for CD3OD, ln for C6D6, lw for D2O, lh for CD2Cl2) |
| lb # | input the line broadening, if value not given (e.g. only lb entered) user is prompted |
| li | list integrals; output can be set to the screen, else output is printed |
| lipp | list integrals with peaks; output can be set to the screen, else output is printed |
| lock | to lock on a deuterium lock signal, the solvent can be specified or selected from the listing that pops up |
| multizg # | run # experiments consecutively (if no number input, you are prompted for one) |
| m# | run multizg for # experiments (# = 20 for 210 experiments) |
| n | (nxtp) next point in processing (e.g. T1 data) |
| new | (edc) call up an existing data set or create a new data set |
| o1 # | input the frequency of the applied pulse for the observed nucleus (centre of the spectrum), if value not given (e.g. only o1 entered) user is prompted, default units is Hz, use o1p for ppm; o2 and o3 for additional nuclei |
| p | (apk) perform automatic phase correction of the spectrum |
| paropt | run the parameter optimisation routine |
| pk | phase correction using set values |
| popt (or POPT) | do parameter optimisation in a more complex way |
| ppg | display the pulse program pictorially |
| pps | list peaks on screen |
| q | quick run, td 64 k, D1 0, a, gs (e.g. proton spectrum) |
| r | return |
| rb | list recent shim files for the BBO probe (equivalent to: rsh BBO*10*pi) |
| ri | list recent shim files for the BBI probe (equivalent to: rsh BBI*10*pi) |
| rk | list parameter files for KDK (equivalent to: rpar kdk*) |
| re # | move to the experiment number # , or if a experiment name has been supplied, to that experimental data set (must already exist, if not, create with e1) |
| rep # | move to processing # of that fid (must already exist, if not, create with e1) |
| rga | automatically set the receiver gain |
| ro | turns on/off the spinner, prompted for yes/no but variables (y/n) not taken, does not work in a macro |
| rotof | turns off the spinner (macro to call tune file of same name) |
| roton | turns on the spinner (macro to call tune file of same name) |
| ROTATION ON/OFF | turns on/off the spinner, only works when given in a tune file and letters capitalised |
| rpar | read the acquisition parameters from a file (names & wildcards can be used) |
| rp | list parameter files for bio (equivalent to: rpar bio*) |
| rser # | retrieve the first FID and move it to ~TEMP 1 file |
| rsh | read the shim values from a file (names & wildcards can be used) |
| rsk | list shim files for KDK (equivalent to: rsh kdk*) |
| rq | run a quick spectrum (equivalent to: td 128 k, D1 0, ns 1, g) |
| r# | read the acquisition parameters from a temporary file saved as w# (# = 10) |
| s | initiate the search subroutine |
| sab | apply baseline correction using user-defined points selected in basl |
| setres | open the window to correct program bugs (e.g. if spectrum is frozen, wobb signal is not being displayed, or if commands from the command line are not acted upon), the apply button is then clicked |
| si # | input the adjusted number of points (if equal to td then no zero-filling is performed), if value not given (e.g. only si entered) user is prompted; typically given directly in k allotments (e.g. 1k, 32k, etc.) |
| sino | calculate the signal-to-noise ratio for the displayed peak |
| sleep(#) | wait for # seconds, function used in au programs |
| splitser | split series file into separate 1D files |
| spoffs# | input data for SPOFF# variable (note the additional “s” in the command, no spacing, & change of case), can enter input on line as “spoffs# $”or wait for the prompt, value can only be given in Hz |
| sr # | set reference correction in Hz relative to Bruker internal scale; if value not given, e.g. only sr entered, user is prompted |
| ss | input data for SPOFF2 variable ( spoffs2) |
| stop | stops the acquisition but does not save the data |
| sw # | input the sweep width of the spectrum, if value not given (e.g. only sw entered) user is prompted; default units is ppm, use swh for Hz |
| t | (setti) edit the title |
| t1 | move to the t1/t2 processing area |
| td # | input the number of sampling points, if value not given (e.g. only td entered) user is prompted, normally used to adjust the acquired digital resolution, typically given directly in k allotments (e.g. 1k, 32k, etc.) |
| teset ### | set temperature to ### (in K) and go there (redundant to use te to set temperature and then teset to change it) |
| touser | to override permission stop on accessing data, parameters |
| tr | transfer the data collected thus far during an accumulation for processing |
| ts | start autoshim for z1 (2), z2 (3), z3 (4), x (2), xz (4), y (2), yz (4) {macro to call tune file “standard”} |
| tsoft | start autoshim for z1 (1), z2 (1), z3 (2) {macro to call tune file “softz”} |
| tss | start autoshim for z1 (2), z2 (3), z3 (4) {macro to call tune file “standardz”} |
| tune | start tune (autoshim) file, prompted with list if name not supplied (wildcards can be used) |
| tune auton/autof | turns on/off the autoshim, works by running the tune file auton or auton |
| tune roton/rotof | turns on/off the spinner, works by running the tune file roton or rotof |
| vish | see shim parameter values |
| w | (wobb + a) start wobbing (for tuning and matching) and move to the acquisition window |
| WAIT | wait command (time ?), only works in a tune file and letters capitalised |
| wpar | save the acquisition parameters to file (names & wildcards can be used) |
| wsh | save the shim values to file (names & wildcards can be used) |
| w# | save the acquisition parameters to a temporary file (# = 10) |
| xfb | process 2D parameter set |
| xf2 | Fourier transform in the f2 dimension only (xf1 for f1 dimension) |
| xf2p | apply phase correction in the f2 dimension only with parameters in file (xf1p for f1 dimension) |
| xt | (expt) request the total time that the experiment requires |
| xw | (xwp) go to xwinplot plotting function |
| zg | (zero go) zero the data and start a new acquisition (the old data is overwritten) |
| 0b | Fourier transformation with 0 Hz line broadening applied (lb = 0 + efp) |
| 1 * # | adjust parameter * for the f1 dimension, if appropriate value not given (e.g. only * entered) user is prompted; e.g. enter “1 sr” for reference correction of the f1 dimension |
| 1b | Fourier transformation with 1 Hz line broadening applied (lb = 1 + efp) |
| 1r | move to processing area 1 (equivalent to rep 1) |
| 1s ns | give the total number of scans acquired (blocks  scans/block), 1s can also be used for other parameters |
| 2 * # | adjust parameter * for the f2 dimension, if appropriate value not given (e.g. only * entered) user is prompted; e.g. enter “2 sr” for reference correction of the f2 dimension; n.b. not entering 2 accesses the observed nucleus and thus provides the same result, hence the use of 2 is redundant |
| 2b | Fourier transformation with 2 Hz line broadening applied (lb = 2 + efp) |
| 2r | move to processing area 2 (equivalent to rep 2) |
| 3b | Fourier transformation with 0.3 Hz line broadening applied (lb = 0.3 + efp) |
| 4b | Fourier transformation with 0.1 Hz line broadening applied (lb = 0.1 + efp) |
| 5b | Fourier transformation with 5 Hz line broadening applied (lb = 5 + efp) |
| 9r | move to processing area 999 (equivalent to rep 999) |